AUTOMATIC SCORING OF DICENTRIC CHROMOSOMES

Dicentric scoring consists of a sequence of 4 relatively independent

stages, (i) metaphase finding, (ii) digitisation of metaphases at

high resolution, (iii) dicentric finding, and (iv) review of the

putative dicentric chromosomes found by the machine.  The events judged

to be true dicentrics may then be used either to calibrate a

dose-response curve, or in the case of an unknown sample, used as input

to a yield-dose curve to obtain the dose estimate.  The 4 stages are

outlined in the following sections, with a review of the performance

achieved and prospects for further improvement.

METAPHASE FINDING

For this trial, the metaphase finding hardware (below) and software

marketed by Image Recognition Systems on their Cytoscan machine and

developed originally by MRC was used effectively unmodified [R1, R2, R3].

The performance on blood lymphocyte metaphases is very good.  However in

a machine devoted almost exclusively to aberration scoring for dosimetry,

the Cytoscan metaphase finder is considerably faster than necessary and

it is likely that a simpler and slower implementation would be more

suitable (some time estimates of the various stages are given below).

AUTOMATIC DIGITISATION

After metaphase finding, the microscope objective must be changed to

x100, oil applied to the slide and the optics set up for optimal image

quality, all manually.  Alternative strategies that might be

investigated are outlined below.  The list of metaphase found and ranked

in quality order in the metaphase search phase is retrieved from disc

and the cells relocated in order.  Each cell is focussed onto the

imaging array by a focus search procedure (see below), the image is

captured, thresholded and background removed in order to compress the

data by a factor of (typically) 10, and the metaphase image stored on

disc [R4].  The cycle of relocation, focus, capture, storage is repeated

automatically until the entire list of metaphases has been captured.

Early in the trial period it became clear that the original auto-focus

search technique for high magnification image capture [R4]

could not adequately cope with focus (Z-axis) motion

backlash on a commercial Cytoscan.  The problem arises in the focus

motion gear box, and although replacing plastic by metal gears improved

matters, an alternative search strategy was developed.  In the original

strategy, a focus function was developed by interpolation from a search

through a series of quite widely spaced steps.  The stage was then moved

to the predicted optimal Z position.  Because of backlash this proved

inadequate, and the new strategy consists of three passes.  The first is

as before, finding a focus function.  In the second pass, the stage is

moved from below the optimal setting in single (0.1 micron) steps until

the autofocus hardware [R1, R3, R4] indicates that the optimal position

has been passed.  The stage is then moved back to the start of the

second search pattern and moved until the point of maximum focus measure

achived in the second pass is again obtained; at this point the

metaphase cell is assumed to be in focus.  This new technique has

resulted in dependably good auto-focus, at the cost of a considerable

increase in the time taken by this phase of the digitisation procedure.

Typically, more than 95% of cells found by the metaphase finder are

digitised in correct focus; the time taken is approximately 40 seconds

per cell.

DICENTRIC FINDING

Following digitisation, the system can be made to continue automatically

with the analysis of the digitised metaphase for dicentric chromosomes.

This stage is implemented by a single program but consists of a sequence

of two relatively independent stages, of segmentation and centromere

classification respectively.

Segmentation

The digitised data has been thresholded and the connected components

predominantly represent single chromosomes.  However, there are

frequently image regions that are not chromosomes; the most common other

types are interphase nuclei, small areas of stain debris, small areas of

stained cytoplasm, and clusters of touching or overlapping chromosomes.

The task of the segmentation program is to recognise these

non-chromosomal regions and to deal with them appropriately; i.e. to

reject nuclei, debris and cytoplasm, and to separate clusters into

individual chromosomes.  So far as we are aware, the system implemented

during the trial period is the only example of a fully automatic

chromosome segmentation system.

Performance of the automatic segmentation has been measured

on a set of about 450 cells.  Overall, about 90% of clusters

are correctly detected and segmented, while approximately

1 chromosome per cell is misclassified as a cluster.

(Check exact figures with Liang !!).

Centromere classification

Image regions accepted as chromosomes after automatic segmentation are

analysed in order to count the centromeres.  The centromere classifier

is essentially as described in [R5] with some incremental improvements.

Briefly, a set of candidate centromeres is found by a first stage

classifier that is "optimistic".  Knowledge that almost all chromosomes

have just one centromere permits the selection of a subset of the

candidates for which it can be assumed that "almost all" are in fact

true centromeres.  The resulting partition of the set of candidates is

used to obtain a classifier that is adapted to the particular conditions

of the cell.  The candidate set is re-classified by this second stage

classifier.

Performance of the classifier is measured in terms of the proportion of

true positive dicentrics detected, and the number of false positives.

The former depends on the radiation exposure of the sample, while the

latter is expected to be almost independent of dose and depends solely

on the number of cells analysed.  It was clearly impractical to score

many thousands of cells entirely visually in parallel with machine

scoring, and so the true positive detection rate has been estimated from

three sources of evidence:

(1) A pilot set of slides of X-irradiated blood which were scored

completely by eye, and then by machine.

(2) A set of 8 slides drawn from the calibration set which were also

scored by eye and by machine (see below).

(3) Comparison of the achieved dose-response with published data for

Co-60 gamma irradiation [R6].

Each of these experiments suggested that the underlying detection rate

taking account of all sources of error is in the region of 40% - 45% of

the dicentrics actually present in the cells successfully analysed by

the system.  This compares with 48% in the original MRC data set.

It should be borne in mind that there is a trade-off between true and

false positive rates.

During the course of the trial it became clear that the performance of

the centromere classifier did not match up to the predictions made on

the basis of a quite substantial data set of irradiated cells collected

by MRC during 1986-88.  While it has proved possible to obtain a similar

performance in terms of the proportion of true dicentrics detected by

the system, the number of false positives is about three times that

obtained with the earlier data.  In separate sections below, we speculate

as to possible reasons for the high false positive rate and outline some

strategies that should reduce it in the future.

REVIEW OF PUTATIVE DICENTRICS DETECTED BY THE SYSTEM

"Dicentrics" found by automatic analysis are stored on disc for

subsequent review.  The review program has three levels of presentation

of such objects.  By default, the object is presented rapidly (<1 sec)

on the display, together with a "histogram-equalised" image.  The

majority of false positives may be rejected immediately without using

the other presentation methods.  If more visual information is required,

the second level is to present an image of the entire metaphase with the

"dicentric" delineated by a circumscribing box.  This display takes

about 5-10 sec on a Cytoscan (but see below for details of equipment

where it takes 1-2 sec).  The third level of presentation is to relocate

the cell under the microscope; this takes 1-2 sec, but there is no means

of pointing to the particular chromosome in the microscope field.

These options and the operators decision are controlled by a menu

interaction, which include the possibility of "backing up" if a decsion

is taken too quickly and the operator wishes to retract it.  The

totals of true and false positives and of the number of cells which have

successfully passed all analysis stages are recorded.

REJECTING UNSUITABLE CELLS

Not all cells found by the metaphase finder are suitable for analysis.

Taking a pragmatic view, a cell is unsuitable if it's condition tends to

prevent the detection of true positives or leads to a disproportionate

number of false positives.  It is also desirable that cells are

approximately complete for otherwise there will be bias in the

dose-response.  Some cell rejection rules have been incorporated in

the system:

(1) Following digitisation, an object count rejects cells with very few

or very many objects separated by threshold.

(2) Following automatic segmentation, a similar rule is applied.

(3) the mean aspect ratio of objects following segmentation, measured as

the square of perimeter divided by area, is computed and cells

rejected for which it is too high (prometaphase cells and anaphase

cells where the majority of chromatids are widely separated), or too

low (extremely condensed metaphase cells).

(4) The mean number of significant minima in the density profiles of the

chromosomes is computed.  Cells with a high value probably show

partial banding and are rejected since the programs are only

intended to cope with homogeneously stained material.

A more systematic cell selection process is planned (see below).

ASPECTS OF TWO DIFFERENT PREPARATION METHODS

The calibration set consisted of irradiated blood separated into two

portions, one of which was cultured immediately, the other stored for

48 hours.  Early in the calibration stage of the trial it became clear

that there was a visible difference of "quality" between the two.

There were also two measurable differences, firstly a lower true

positive frequency and secondly a higher false positive rate.

Because the same original blood was used it was at first feared that

there was a significant preparation-dependent performance bias in the

analysis algorithms.  A small blind trial was performed to test this

hypothesis: 4 slides from each preparation set, 2 each at 1.5Gy and

2.0Gy, which had not yet been analysed by machine, were scored visually

by competetent cytogeneticists at MRC.  The result was that there was

indeed a very substantial difference in the true proportion of

dicentrics between the two preparation methods; the reason is

unexplained.  It is believed that the difference in the false positive

rate reflects the difference in the visual "quality"; this would not be

unexpected (and the quality difference itself is also expected).

TIMING DATA

Metaphase finding takes about 5 minutes per slide or approximately

1 second permetaphase.

Automatic digitisation takes about 40 seconds per metaphase.

Segmentation and dicentric scoring takes about 2 minutes

per metaphase, although this figure is reduced considerably

if display functions are disabled.

Thus overall the throughput of the system before review is

about 20 cells per hour.

Review is of course unpredictable, since a cytogeneticist has

the choice of options outlined above and may wish to look at

some cells in considerable detail.  However in low dose specimens

it is certainly possible to review and reject the false positives

at a rate of one per 10 seconds overall, possibly much faster.

Even with the rather high false positive rate of 0.5 per cell

measured on some of the calibration data, this implies an

overall throughput in review of 720 analysed cells per hour.

It is clear therefore that improvements in computation speed of the

automatic analysis is an important future development (see below).

CYTOSCAN/FIP HARDWARE

The machines consist of a Nikon or Zeiss microscope with a computer

controlled motorised scanning stage capable of 1 micron steps in the

Y axis and 0.1 micron steps in X and Z.  The scanner is a 1024 element

linear charge coupled photodiode array, with subsidiary arrays for

auto-focus.  Special hardware pre-processes data from the array, and

software analysis, stage control, and interface to the user are provided

by a Plessey Miproc computer and a multiprocessor based on M68000 and

VME-bus.  The only modofocation to standard Cytoscan hardware was to

provide the facility to measure microscope light level and control the

lamp intensity in a software servo loop.  Except for metaphase finding,

the software components are substantaially different to Cytoscan.

UP-COMING DEVELOPMENTS

The calibration stage of the trial has provided MRC with useful

experience of its software on a data set that is about 15 times larger

than all data used previously for automatic dicentric detection

research.  Two related aspects stand out as being in need of

improvement,

(1) Reduction in the false positive rate.  While it is certainly

possible to reject false positives extremely rapidly at the

review stage, and thus the entire system is already economical,

it is not psychologically appropriate to have so many objects

presented that are clearly not dicentric - the computer achieves

a reputation of "stupidity".  Work is under way with an alternative

centromere classifier.

(2) Many false positives come from cells that are clearly inappropriate

for analysis, in which it would not be possible reliably to find

a dicentric were it to be present.  We intend to treat cell

selection as a pattern recognition problem, training the system

with (i) cells that have false positives and/or false negatives

but no true positive (undesirable) and (ii) cells where either a true

positive or nothing at all has been found.

SOME POSSIBLE ALTERNATIVE STRATEGIES AND HARDWARE

We have investigated replacing the multi-processor M68000 part of the

system with a computer based on the architecturally similar M68030.

The effect is dramatic; digitisation time falls by a few seconds but

analysis times falls by a factor of about six.  We estimate that this

single change can increase throughput of the digitisation and analysis

cycles by a factor of three, and if digitisation and analysis were to

be overlapped then a factor of five can probably be achieved, with a

resulting throughput of about 100 cells per hour.

REFERENCES

R1. Shippey GA, Bayley RJH, Farrow ASJ, Rutovitz D, Tucker JH: ,

"A fast interval processor".

Pattern Recognition 14:345-356, 1981.

R2. Shippey G, Carothers AD, Gordon J,

"Operation and performance of an automatic metaphase finder

based on the MRC Fast Interval Processor".

J Histochem Cytochem 34:1245-1252, 1986.

R3. Lloyd D, Piper J, Rutovitz D, Shippey G,

"A multiprocessing interval processor for automated cytogenetics".

Applied Optics 26:3356-3366, 1987

R4. Stark M, Farrow S, McKie M, Rutovitz D,

"Automatic highresolution digitisation of metaphase cells

for aberration scoring and karyotyping".

In: Automation of Cytogenetics (eds: Lundsteen C, Piper J),

Springer-Verlag, Heidelberg, 1989, 31-43.

R5. Piper J, Towers S, Gordon J, Ireland J, McDougall D,

"Hypothesis combination and context sensitive classification for

chromosome aberration scoring".

In: Pattern Recognition and Artificial Intelligence

(eds: Gelsema ES, Kanal LN),

Elsevier, Amsterdam, 1988, 449-460.

R6. Lloyd.  Co-60 gamma dose-response